ADE7912, ADE7913, ADE7932, ADE7933 Data Sheet Detailed Change List

Parameter specification changes are listed as follows:

- ADC Offset Drift Over Temperature 

From:  	-500 PPM/°C minimum, +500 PPM/°C maximum for V1 channel only

To: 	+200 PPM/°C typical for V1 channel only

- Gain Drift Over Temperature 

From:  	-135 PPM/°C minimum, +135 PPM/°C maximum for current channel
 	-65 PPM/°C minimum, +65 PPM/°C maximum for V1 and V2 channel

To:	-135 PPM/°C minimum, +135 PPM/°C maximum for current channel
 	-85 PPM/°C minimum, +85 PPM/°C maximum for V1 and V2 channel

- Input Current, Iin

From:	 15 nA maximum

To:	 0.015 uA typical, 1uA maximum

This product has achieved UL, CSA and VDE certification and absolute maximum rating changes are listed as follows:

- Regulatory Approvals
From:

| UL | CSA | VDE |
| --- | --- | --- |
| Recognized Under UL 1577 Component Recognition Program | Approved under CSA Component Acceptance Notice 5A | Certified according to DIN V VDE 0884-10 (VDE V 0884-10):2006-12 |
| Single Protection, 5000 V rms Isolation Voltage | Basic insulation per IEC 61010-1, 400 V rms (564 V peak) maximum working voltage | Reinforced insulation, 846 V peak |

Note:  In accordance with UL 1577, each ADE7912/ADE7913 is proof tested by applying an insulation test voltage ≥ 6000 V rms for 1 second (current leakage detection limit = 10 μA).  

To:

| UL | CSA | VDE |
| --- | --- | --- |
| Recognized Under UL 1577 Component Recognition Program | Approved under CSA Component Acceptance Notice 5A | Certified according to DIN VDE V 0884-10 (VDE V 0884-10):2006-12 |
| Single Protection, 5000 V rms Isolation Voltage | Basic insulation per CSA 60950-1-07+A1+A2 and IEC 60950-1 2nd Ed.+A1+A2: - 830 V rms (1173 V peak) maximum working voltage Basic insulation per CSA 61010-1-12 and IEC 61010-1 3rd Ed. (pollution degree 2, material group III, overvoltage category II, III and IV): - 300 V rms (424 V peak) maximum working voltage. Reinforced insulation per CSA 60950-1-07+A1+A2 and IEC 60950-1 2nd Ed.+A1+A2: - 415 V rms (586 V peak) maximum working voltage Reinforced insulation per CSA 61010-1-12 and IEC 61010-1 3rd Ed. (pollution degree 2, material group III, overvoltage category II, and III): - 300 V rms (424 V peak) maximum working voltage. | Reinforced insulation, 846 V peak |
| FILE E214100 | FILE 2758945 | FILE 2471900-4880-0001 |

In accordance with UL 1577, each ADE7912/ADE7913 is proof tested by applying an insulation test voltage ≥ 6000 V rms for 1 second (current leakage detection limit = 15 μA).

- Critical Safety Related Dimensions and Material Properties
From:

| Parameter | Symbol | Value | Unit | Test Conditions/Comments |
| --- | --- | --- | --- | --- |
| Tracking Resistance (Comparative Tracking Index) | CTI | >600 | V | IEC 60112 |

To:

| Parameter | Symbol | Value | Unit | Test Conditions/Comments |
| --- | --- | --- | --- | --- |
| Tracking Resistance (Comparative Tracking Index) | CTI | 400 | V | IEC 60112 |

- VDE Characteristics
From:

| Description | Test Conditions/Comments | Symbol | Characteristic | Unit |
| --- | --- | --- | --- | --- |
| Climatic Classification |  |  | 40/105/21 |  |
| Surge Isolation Voltage | VPEAK = 10 kV, 1.2 µs rise time, 50 µs, 50% fall time | VIOSM | 6000 | V peak |

To:

| Description | Test Conditions/Comments | Symbol | Characteristic | Unit |
| --- | --- | --- | --- | --- |
| Climatic Classification |  |  | 40/085/21 |  |
| Surge Isolation Voltage | VPEAK = 10 kV, 1.2 µs rise time, 50 µs, 50% fall time | VIOSM | 6250 | V peak |

- Thermal Derating Curve
From:

To:

5.   Maximum Continuous Working Voltage Supporting a 50-Year Minimum Lifetime 

     From:  	AC Voltage, Bipolar Waveform   564V peak
                  	DC Voltage, Basic Information   600V peak

     To:  		AC Voltage, Bipolar Waveform   400Vrms
                  	DC Voltage, Basic Information   1173V peak