Table XIX. Pseudo-Code for Initiating a Self-Calibration after Power-On/Reset

Write 03 Hex to Serial Port1	/* Writes to Communications Register Setting Next Operation as Write to Filter Register*/

Write 1332 Hex to Serial Port1	/* Writes to Filter Register Setting a 1 kHz Output Rate in nonCHOP Mode*/

Write 02 Hex to Serial Port	/* Writes to Communications Register Setting Next Operation as Write to Mode Register*/

Write B174 Hex to Serial Port	/* Writes to Mode Register Initiating Internal Full-Scale Calibration for 0 V to +1.28 V Input Range on Channel Pair AIN1/AIN2*/

Wait for RDY Low	/* Wait for RDY pin to go low to indicate end of calibration cycle*/

Write 02 Hex to Serial Port	/* Writes to Communications Register Setting Next Operation as Write to Mode Register*/

Write 9174 Hex to Serial Port	/* Writes to Mode Register Initiating Internal Zero-Scale Calibration for 0 V to +1.28 V Input Range*/

Wait for RDY Low	/* Wait for RDY pin to go low to indicate end of calibration cycle*/

/* The part has now completed self-calibration and is in idle mode*/

1This operation is not necessary if the default values of the Filter Register are the values used in the application.