Document No.	:	18-046277-01 Rev C

Title			:	AD916x-FMC-EBZ Customer Evaluation Board Test Procedure

| REVISION HISTORY |  |  |  |  |
| --- | --- | --- | --- | --- |
| Revision | ECR # | Description of Change | Date | Author |
| C | ECR-082319 | - | 1/31/18 | Saber Bahranifard |

| Required Approvers |  |
| --- | --- |
| Approver Roles | Approver Names |
| Apps Engineer | Saber Bahranifard |

AD916x-FMCC-EBZ Test Procedure – Carl Scire - 03/31/2017

Software Setup

- Install the latest DPG Downlaoder from DAC Software Suite DVD.

- Install ADI FMC EEPROM Generator, https://code.analog.com/gf/project/fmc_eeprom/frs/.

Hardware Setup

- Set SML Signal Generator to 3 GHz 3dBm amplitude and connect it to J31.

- Set a second SML Signal Generator to 130 MHz 3dBm amplitude and connect it to J61.

- Connect Melabs U2 Programmer to XP1 and Mini USB to XP2.

- Connect the signal analyzer to the SMA connector, J32. The spectrum analyzer can be configured with Start Frequency = 50 MHz, Stop Frequency 5 GHz, and Resolution Bandwidth of 300 kHz. You can use an Average/RMS detector setting, if you prefer, and choose Input Attenuation to be 6 dB.

Program

- Open DPG Downloader (Start > Programs > Analog Devices > DPG > DPGDownloader). Wait a few seconds until the power LED on the evaluation board is lit.

- Connect AD9161 (1/2/3/4)-FMCC-EBZ Board to ADS7-V2 using FMC P1 connector, Apply power to ADS7-V2.

- Start->Analog Devices ->ADI FMC EEPROM Generator, click Open File and select the correct EEPROM file for the board you are testing

- Click Read/Write EEPROM and Click write. When finished, press OK.

- Power down ADS7-V2.  Wait until DPG Downloader shows “No devices found”.

- Power on ADS7-V2.

- Open Melabs U2 Programmer and select File – Open, opening the correct .hex SPI Firmware for the AD9161-FMCC-EBZ board you are testing.

- Press program button.  When finished, press OK.

- Open the ACE software, which is located in Start->Programs->Analog Devices->ACE->ACE.

- Run ACE Software and click on AD9161-FMCC-EBZ for testing with AD916X Startup Wizard.

Settings in ACE Using an External Clock Source

- Set front output of a Signal Generator to 3 GHz with 3dBm amplitude and connect it to J31.

- Set a separate Signal Generator to 130 MHZ, 3dBm amplitude and connect it to J61.

- Make sure jumper shunt from JP1 is removed.

- Click on AD9161-FMCC-EBZ in ACE.

- In AD916X Startup Wizard Summary - Select External Direct Clock (J31), set Reference Clock (Osc or Ext): 130MHz, FDAC: 3GHZ.  As seen below and click “Apply”.  Wait until “Apply” changes to “Modify”.

15.In AD916X Startup Wizard Summary - Select External Direct Clock (J31), set Reference Clock (Osc or Ext): 130MHz, FDAC: 3GHZ.  As seen below and click “Apply”.  Wait until “Apply” changes to “Modify

- In DPG Downloader, click on “Add Generated Waveform”, and then “Single Tone”. A single tone panel will be added to the vector list. Enter the sample rate, or DAC clock frequency in this case, 1.5 GHz.    Default “Resolution” should be 16 bits. Set Desired Frequency to 750 MHz.  Keep “Amplitude” as 0dB. Uncheck the “Unsigned Data” check box. Set up the DPG Downloader with the right configuration for testing the AD9161-FMCC-EBZ.  Select Generate Complex Data by placing a check mark at Generate Complex Data.  Set Interpolation Mode at >1X.

- In the lower portion of the screen, select 1I: Single Tone - 750.092 MHz; 0.0 dB; 0.0o (In-Phase) as the Data Vector.  See below.

- In the lower portion of the screen, select 1Q: Single Tone - 750.092 MHz; 0.0 dB; 0.0o (Quadrature) as the Data Vector.  See below

### External Clock Hardware Setup with only 1 Signal Source

- Set front output of Signal Generator to 3 GHz with 3dBm amplitude and connect it to J31.

- Use 10 MHz reference output from the rear of Signal Generator and connect it to J61.

- Make sure jumper shunt from JP1 is removed.

- Run ACE or clear ACE from last session by going to File - Close Session.

- Jumper shunt should not be installed at JP1.

- In ACE AD916X Startup Wizard Summary- Select External Direct Clock (J31), set Reference Clock (Osc or Ext): 10MHz, FDAC: 3GHZ.

- Go directly to DPG Downloader

In DPG Downloader set Data Rate to 1.5 GHz.  Default “Resolution” should be 16 bits. Set Desired Frequency to 10 MHz.  Keep “Amplitude” as 0dB. “Unsigned Data” box should be unchecked.

In the lower portion of the screen, Select Generate Complex Data by placing a check mark at Generate Complex Data.  Set Interpolation Mode at >1X.

- In the lower portion of the screen, select 1I: Single Tone – 9.979 MHz; 0.0 dB; 0.0o (In-Phase) as the Data Vector.  See below.

- In the lower portion of the screen, select 1Q: Single Tone – 9.979 MHz; 0.0 dB; 0.0o (Quadrature) as the Data Vector.  See below

- Click the  button to download the pattern from the computer to the ADS7 unit, wait for the Play () button to become active, and then click the Play ( ) button to begin vector playback to the AD9161. The “SYNC Status” should show as the green check mark.

- Approximately a 10 MHz single tone output should now also appear on the signal analyzer as seen on the screen shot below

- Shut down ADS7-V2, remove cable from J31 and install jumper shunt at JP1.

Testing AD9161-FMCC-EBZ using On Board Clock

Set the Signal Generator to 120 MHz with 3dBm amplitude and connect it to J61.

In DPG Downloader set Data Rate to 2.5 GHz.  Default “Resolution” should be 16 bits. Set Desired Frequency to 1.25 GHz.  Keep “Amplitude” as 0dB. “Unsigned Data” box should be unchecked.

In the lower portion of the screen, Select Generate Complex Data by placing a check mark at Generate Complex Data.  Set Interpolation Mode at >1X.

- In the lower portion of the screen, select 1I: Single Tone – 1.250 GHz; 0.0 dB; 0.0o (In-Phase) as the Data Vector.

- In the lower portion of the screen, select 1Q: Single Tone – 1.250 GHz; 0.0 dB; 0.0o (Quadrature) as the Data Vector.

- Clear ACE from last session by going to File - Close Session

- Jumper shunt should be installed at JP1.

- Restart ADS7-V2.

- Default is AD9161-FMCC-EBZ for testing with on board ADF4355 Clock + ADCLK914.

- Use default settings and hit “Apply”.

- Click the  button to download the pattern from the computer to the ADS7 unit, wait for the Play () button to become active, and then click the Play ( ) button to begin vector playback to the AD9161. The “SYNC Status” should show as the green check mark.  See below.

Approximately an 800 MHz single tone output should now also appear on the signal analyzer as seen on the screen shot below.