Document No.	:	18-073531-01 Rev A

Title			:	ADL8100-EVALZ Customer Evaluation Board Test Procedure

| REVISION HISTORY |  |  |  |  |
| --- | --- | --- | --- | --- |
| Revision | ECR # | Description of Change | Date | Author |
| D | ECR-114068 | Initial Release | 04/24/2023 | Love Syl Venenoso |

| Required Approvers |  |
| --- | --- |
| Approver Roles | Approver Names |
| Product Engineer | Love Syl Venenoso |

| Model Name | C-Spec | Tester Program |
| --- | --- | --- |
| ADL8100-EVALZ | C-11095 A | ADL8100-EVALZ.par |

Test 1: S-Parameter Sweep

MARKI BTN20050 BIAS TEE

MARKI BTN 20050 DC BLOCK

VDD (PS1)

Network Analyzer

Note: Not drawn to scale.

VBIAS

GND

Equipment Needed:

1 Power Supply – Keithley 2400 or its equivalent

1 2-port Network Analyzer – N5245B or its equivalent

1 Marki BTN20050 DC BLOCK

1 Marki BTN20050 Bias Tee

Setup:

Connect the Network Analyzer as shown

Port 1 – RFIN

Port 2 – RFOUT

Network Analyzer Settings:

Min/Max Frequency: 10 MHz to 22 GHz

Number of Points: 501 points

Input Power: -25 dBm

Setting up notes:

- Calibrate the Network Analyzer to compensate for the loss of both input and output cables.

- Record the S11, S12, S21(Gain), S22 of the DUT. Compare the data with the plot below.

Biasing steps:

- Turn on VDD (PS1) = +5 V, typical IDQ = 220 mA

- Apply the RF signal.

- Log the data.

Power off:

- Turn off the RF signal.

- Turn off VDD.

S21, S11 and S22 plot

Test 2: ATE Setup

Equipment and Accessories Needed:

1 Power Supply – Keithley 2400 or its equivalent

1 Signal Generator – Keysight E8267D or its equivalent

1 External Amplifier – N4985A or its equivalent

1 Power meter – Anritsu ML2496A or its equivalent

1 Power Sensor – Anritsu MA2445D or its equivalent

1 20 dB pad

1 Marki BTN20050 DC Block

1 Marki BTN20050 Bias Tee

Connect the setup as shown below.

RFIN

RFOUT

Computer

Power Meter

Amplifier

20 dB pad

Power Sensor

RF Source

VDD (PS1)

Note: Not drawn to scale.

MARKI BTN20050 DC BLOCK

MARKI BTN20050 BIAS TEE

VBIAS

GND

Test Parameters: