Document No.	:  18-069659-01-A

Title			:  EVAL-AD4030-24FMCZ Customer Evaluation Board Test Procedure

| REVISION HISTORY |  |  |  |  |
| --- | --- | --- | --- | --- |
| Revision | ECR # | Description of Change | Date | Author |
| A |  | Initial Release | 11/16/2021 | Clarence Mayott |
| A |  | Corrected typos and removed test steps that refer to channel B | 11/18/2021 | Jim Catt |

| Required Approvers |  |
| --- | --- |
| Approver Roles | Approver Names |
| Apps Engineer | Clarence Mayott |
| Applications Manager | Jim Catt |

# Equipment List

The following equipment is required to execute the test procedure describe here. All equipment must have a current calibration certification.

Table 1. Test equipment list

| Equipment name | Model number/part number | Required Options | Description |
| --- | --- | --- | --- |
| DVM | Fluke Model 87 or equivalent | N/A | 4.5 digit digital volt-meter |
| Sine Wave Source | Stanford Research SR1 or equivalent like AP SYS-2722 | Ultra-Low Distortion with XLR male output. Specs: THD+N < -112dB at 1kHz and at 4Vrms | Low distortion sinewave generator. Used as input stimulus for the DUT |
| USB to SD Card adapter | Insignia or equivalent | https://www.amazon.com/Insignia-Advanced-Memory-Card-Reader/dp/B074M7DYRV | USB based SD Card Writer/Reader to program the SD card. |
| XLR to SMA filter adapter board |  | Provided by ADI | The XLR connector connects to the output of SR1 and the SMA connector connects to the AD4030-24 EVAL board |
| 2-SMA (male) cables |  | General Lab accessory | Any length is fine but preferably matched. Connects from the output of XLR to SMA filter adapter board to the SMA inputs of the AD4030-24 EVAL board. |
| PC | OS version Windows 7 or newer | USB port | For performing EVAL board tests and for programming SD cards. |
| Zedboard™ | Zedboard™ | Golden Zedboard provided by ADI | Use the provided Zedboard™ for capturing Data |
| Pre-programmed SD Card | 16GB SanDisk | Provided by ADI | Used to setup Zedboard™ |
| USB Pen Drive |  | Any brand will work. | To setup SR1 for generating sine wave |
| SDP-H1 | SDP-H1 | Buy from this link: https://www.analog.com/en/design-center/evaluation-hardware-and-software/evaluation-boards-kits/eval-sdp-h1.html | Used to program the EEPROM chip on EVAL-AD4030-24FMCZ. Buy from this link: https://www.analog.com/en/design-center/evaluation-hardware-and-software/evaluation-boards-kits/eval-sdp-h1.html |
| 2x USB-A male to micro USB male cable |  | Any brand will work. Purchase link: https://www.digikey.com/en/products/detail/cnc-tech/102-1072-BL-00050/3064848 | One cable gets connected to Zedboard’s™ USB UART port and other cable gets connected to Zedboard’s™ USB OTG port |

Table 2. Software list

| Name | Version | Download link | Description |
| --- | --- | --- | --- |
| EVAL-AD4030-24FMCZ.zip | 1.0 | Use the provided link | Zip file containing all the necessary software packages to test the EVAL-AD4030-24 boards |
| Spyder IDE | 3.7.6 | https://repo.anaconda.com/archive/Anaconda3-2020.02-Windows-x86.exe | Anaconda Spyder IDE for running python program |
| Win32Disk Imager | 1.0.0 | https://sourceforge.net/projects/win32diskimager/files/latest/download | Application used for SD card programming |
| TeraTerm | 4.105 | https://tera-term.en.lo4d.com/download/mirror-hs1 | Application used to verify programmed SD card |
| 7-zip file manager | 19.00 | https://www.7-zip.org/a/7z1900-x64.exe | Application used to extract the zip folders. |
| Compressed SD Card disc image file to program the SD cards (2021-07-28-ADI-Kuiper-full-RC-ad4030-v6.zip) | 1.0 | Use the provided link | Compressed Disc Image file to program the SD cards |

# Introduction

This test procedure consists of multiple parts. These parts address three basic tasks required to prepare the EVAL-AD4030-24 for kitting and shipping.

- Programming the EEPROM on the EVAL-AD4030-24FMCZ

- Conducting the electrical test of the EVAL-AD4030-24FMCZ

- Programming the SD memory card with the boot image that is required for the controller board (ZedBoard).

There are different PC software tools required to accomplish these tasks. The following section addresses the installation of all the tools.

# Part 1 - EVAL-AD4030-24FMCZ software package installation

The software items included in this section address all the tools needed to program the EEPROM, test the board, and program the SD memory cards.

Follow the steps below to create the appropriate folder structure in the PC and unzip the software:-

- Download the software packages in Table 2.

- Create a folder labelled as “projects” in C drive. No need to create this folder if it already exists.

- Create another folder named eval-ad4030_files in C:\projects. Extract the EVAL-AD4030-24FMCZ.zip to C:\projects\eval-ad4030_files.

- Please ensure all the extracted files are located in C:\projects\eval-ad4030_files directory.

- The extracted file should contain the following files:

  - FMC_EEPROM_Setup_1.0.6920.31141.exe,

  - EVAL-AD4030-24FMCZ_production_build.fmceeprom (for SDP-H1),

  - EVAL-AD4030-24_SR1_config.xml (config file for SR1),

  - libiio folder,

  - pyadi-iio folder,

  - sin_params.py,

  - ad4030fmcz_capture_GUI_new_HDL.py (test software for EVAL-AD4030-24FMCZ),

  - PlutoSDR-M2k-USB-drivers.exe,

  - libiio-0.21.g56bf68-Windows-setup.exe,

  - 2021-07-28-ADI-Kuiper-full-RC-ad4030-v6.zip

- Install the ADI FMC EEPROM generator by launching the FMC_EEPROM_Setup_1.0.6920.31141.exe application. Follow the on-screen instructions for the setup.

- Install the IIO device driver for windows by launching the libiio-0.21.g565bf68-Windows-setup.exe application and follow the instructions on the screen.

- Install the Windows USB drivers by launching the PlutoSDR-M2k-USB-Drivers.exe application. Follow the on-screen instructions for the setup.

- Download and install the Anaconda Spyder IDE using the following link: https://repo.anaconda.com/archive/Anaconda3-2020.02-Windows-x86.exe

- Download and install win32diskImager from this link: https://sourceforge.net/projects/win32diskimager/files/latest/download. Wait for the download to begin and once downloaded please launch the win32diskimager-1.0.0-install.exe application and follow on-screen instructions to complete the installation process

- Download and install TeraTerm from this link: https://tera-term.en.lo4d.com/download/mirror-hs1. Wait for the download to begin and once downloaded please launch the teraterm-4.105.exe application and follow on-screen instructions to complete the installation process.

- Download and install the 7-zip file manager from this link: https://www.7-zip.org/a/7z1900-x64.exe application and follow the on-screen instructions to complete the installation process.

# Part 2 - EEPROM Programming on the EVAL-AD4030-24FMCZ

Program the EEPROM chips on all the EVAL-AD4030-24FMCZ boards by using the SDP-H1 board and the ADI FMC EEPROM Generator tool. Follow the below steps to program the EEPROM chip on each EVAL-AD4030-24FMCZ board: -

- Launch the ADI FMC EEPROM Generator tool from this location: C:\Program Files (x86)\Analog Devices\FMC EEPROM Generator.

- This will launch a GUI as shown in Figure 1. Click on Open File button located on Bottom and browse to the directory where the EVAL-AD4030-24FMCZ_production_build.fmceeprom file is located.

Click here

Figure 1. ADI FMC EEPROM Generator Tool

- Connect the SDP-H1 and the EVAL-AD4030-24FMCZ using the FMC connector as shown in Figure 2. Power up the SDP-H1 board using the +12V AC adapter cable and connect the USB cable from J1 connector to the host PC as shown in Figure 2. Please ensure that the “BF_POWER” and the “SYS_PWR” LEDs on SDP-H1 are turned on.

Figure 2. Hardware Setup to program the EEPROM chip

- WRITE EEPROM Step: Click on Read/Write EEPROM button located on the bottom of the GUI. This opens another window labelled as EEPROM as shown in Figure 3. Then click on the Write button located on the bottom of the EEPROM window as shown in Figure 3. This programs the EEPROM.

Figure 3. ADI FMC Generator Tool – Write EEPROM chip

- VERIFY EEPROM Step: Close the ADI FMC EEPROM Generator tool GUI and re-launch the ADI FMC EEPROM Generator tool GUI program. Click on the Read/Write EEPROM button. This opens another window labelled as EEPROM as shown in Figure 3Figure 4. Click on the Read button on the EEPROM window to read back the programmed values and ensure that the readback contents matches with the fields inside the red box as shown in Figure 4.

Figure 4. ADI FMC Generator Tool – Read EEPROM chip’s values

- After verifying the EEPROM contents, close the ADI FMC Generator tool and remove the power supply from the SDP-H1 by disconnecting the +12V Power adapter.

- Repeat Steps 1 to 7 for the remaining EVAL-AD4030-24FMCZ boards.

# Part 3 – EVAL-AD4030-24FMCZ Electrical Test

## 3.1 Board connections and Test Setup for EVAL-AD4030-24FMCZ

Figure 5. EVAL-AD4030-24FMCZ connections and Test Setup

Note: This figure shows the EVAL-AD4630-24FMCZ, which is the two-channel version of the EVAL-AD4030-24FMCZ.

This section outlines the steps required to test a whole batch of EVAL-AD4030-24FMCZ boards.

Required equipment:

- Golden Zedboard™

- Golden, pre-programmed SD card. (If the golden pre-programmed SD card is not available, follow the steps outlined in the Part 4 - SD Card Programming Procedure section to burn a SD card.)

- XLR to SMA low pass filter adapter board/test jig.

- Precision sine wave signal source (Stanford Research SR1, Audio Precision® SYS-2722, which has a XLR male output connector, or equivalent.)

Setup:

- If the SR1 is used for the test, see appendix 1 for instructions on how to configure the signal generator using the EVAL-AD4030-24_SR1_config.xml file that is provided with this procedure.

- To manually setup the SR1, open the Analog Generator window and apply the following settings:

  - Fs = 512kHz

  - Mode = Mono (single tone)

  - Output = Balanced float with differential impedance = 50Ω

  - sine wave frequency = 1kHz

  - sine wave amplitude =3.5Vrms.

  - Check the box in between A/B and Lock as shown in Figure 6 to lock Ch. A with respect to Ch. B.

  - Also check the box below Invt. in Ch. B to ensure Ch.B is inverted with respect to Ch. A as shown in Figure 6

- Please ensure that the Analog Generator window’s setup matches with the settings shown in Figure 6.

Figure 6. Analog Generator window settings on SR1

- Connect the output of SR1 to the XLR to SMA filter board and then connect the output of the XLR to SMA filter board to IN0_A+ and IN0_A- (Channel 0/Channel A) of the EVAL-AD4030-24FMCZ board.

Follow the below steps to finish the remainder of hardware setup in order to test a batch of EVAL-AD4030-24FMCZ boards.

(DO NOT POWER UP THE ZEDBOARD UNTIL INDICATED IN THE PROCEDURE)

- Insert the golden pre-programmed SD card to the bottom of the Zedboard which has a slot for the SD Card as shown in Figure 5.

- Ensure that the MIO jumper settings on the Zedboard matches with the MIO jumper settings shown in Figure 7.

Figure 7. MIO Jumper (JP7-JP11) Settings on Zedboard

- Connect the micro-USB end of the micro-USB to USB cable to the Zedboard’s USB OTG Port and connect the USB end of the micro-USB to USB cable to the host PC’s USB port as shown in Figure 5.

- Connect the micro-USB end of the micro-USB to USB cable to the Zedboard’s USB UART Port and connect the USB end of the micro-USB to USB cable to the host PC’s USB port as shown in Figure 5.

- Connect the 12V adapter to the +12V connector (J20) on the Zedboard as shown in Figure 5.

Setup the Spyder IDE to finish up the software setup:-

- Launch the Spyder IDE from the Start menu.

- When the Spyder IDE opens, setup the library path in python path manager as shown in Figure 8. The paths to be added are:

  - C:\projects\eval-ad4030_files

  - C:\projects\eval-ad4030_files\libiio\bindings\python

  - C:\projects\eval-ad4030_files\pyadi-iio

  - C:\projects\eval-ad4030_files\pyadi-iio\adi

Figure 8. PYTHONPATH manager settings

- Ensure that the Preferences settings match Figure 9.

Figure 9. Setting up Preferences for displaying image in separate window

- Restart the Spyder IDE for the changes made in PythonPath Manager and Preferences to take effect.

- After re-starting the IDE, open the ad4030fmcz_capture_GUI_new_HDL.py program from C:\projects\eval-ad4030_files directory.

## 3.2 Performance testing of the EVAL-AD4030-24FMCZ board

- Connect an EVAL-AD4030-24FMCZ board with the Zedboard™ as shown in Figure 5.

- Power up the Zedboard™ by sliding the POWER switch (SW8) on the Zedboard to the ON position as shown in Figure 17. Ensure that the Green POWER LED (LD13) and the Blue Done LED (LD12) on the Zedboard™ turns on as well as shown in Figure 17.

- Wait for the Red LED (LD7) on the Zedboard™ to start blinking. It may take up to 20 - 25 seconds for the RED LED to start blinking.

- After the Zedboard™ powers up, ensure that the USB OTG connection from Zedboard to host PC is detected in the Device Manager as shown in Figure 10.

Figure 10. Ensure USB OTG connection is detected in Device Manager

- The EVAL-AD4030-24FMCZ board is powered from the Zedboard through the FMC connector. Measure the on-board DC voltages on the EVAL-AD4030-24FMCZ board at the test points listed in Table 3 using a 4.5 digit handheld Fluke 87 multimeter.

- DC Performance test - Verify that the measured voltages are within the specified limits as listed in Table 3.

Table 3. DC Voltage Test Points and Measurement Limits

| Test Point | Min | Max |
| --- | --- | --- |
| VAMP- | -3.353V | -3.287V |
| VAMP+ | 5.366V | 5.474V |
| REFIN | 4.950V | 5.050V |
| +5V | 5.267V | 5.373V |
| +1.8V | 1.774V | 1.809V |
| VIO | 1.832V | 1.869V |
| +3.3V | 3.256V | 3.322V |

- AC Performance Test: Run the Python program by pressing the “F5” key on keyboard or by clicking the green arrow icon on the Spyder IDE. When the program launches the user will see the application GUI (Figure 11). The console in the Spyder IDE (Figure 12) will output useful debug messages, which will help identify the issues in capturing data from the EVAL-AD4030-24FMCZ hardware.

Figure 11. AD4030-24 Evaluation Software Front Panel (Data Capture Plots Tab View)

- Click on the “Establish Connection” button and if the connection is successful, then the console in the Spyder IDE will display a message “Successfully established communication with Zedboard” as shown in Figure 12 .

Figure 12. Spyder IDE Console Output

- Return to the Data Capture Plots tab of the AD4030-24 GUI window and select Channel A from the ADC Channel drop-down list.

- Click on the Single Capture button. The GUI will plot the Time Domain signal along with the FFT of the captured data as shown in Figure 13. Note the values of SNR and THD and ensure that they are within the limits specified in Table 4.

Figure 13. Time Domain plot and the FFT domain plot

- This concludes the electrical test of the board. Turn off the Zedboard by sliding the POWER Switch SW8 to OFF position. Disconnect the EVAL-AD4030-24FMCZ board from the Zedboard™ and replace with another EVAL-AD4030-24FMCZ board.

- Turn on the Zedboard and ensure that the USB OTG connection from Zedboard to host PC is detected in the Device Manager as shown in Figure 10. Repeat steps 1 to 10 to test the remaining EVAL-AD4030-24FMCZ boards.

Table 4. AC Dynamic Test limits (Sine Wave Testing Results limits)

| Channel No. | SNR (dB) |  |  | THD (dB) |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | min | typ | max | min | typ | max |
| A | 104 | 105.23 | 108 |  | -120.65 | -112 |

# Test Termination Procedures

All board tests are complete at this point.

# Part 4 - SD Card Programming Procedure

Please follow the below steps to program a batch of SD cards. For this step a known good Zedboard™ and a known good AD4030-24 EVAL board are required. Please refer to the

Part 3 – EVAL-AD4030-24FMCZ Electrical Test

3.1 Board connections and Test Setup for EVAL-AD4030-24FMCZ section in order to get a good AD4030-24 EVAL board.

(DO NOT POWER UP THE ZEDBOARD UNTIL INDICATED IN THE PROCEDURE)

- Launch the 7-zip file manager and extract the 2021-07-28-ADI-Kuiper-full-RC-ad4030-v6.zip in the same directory i.e C:\projects\eval-ad4030_files as shown in Figure 14.

Figure 14. 7-zip file manager to extract the SD card image file-2021-07-28-ADI-Kuiper-full-RC-ad4030-v6.img

- Use the USB SD card reader shown in this link: https://www.amazon.com/Insignia-Advanced-Memory-Card-Reader/dp/B074M7DYRV

- Insert the SD card to into the SD Card writer/reader and connect to the host PC. Then launch the win32diskImager application as an Administrator and locate the 2021-07-28-ADI-Kuiper-full-RC-ad4030-v6.img file (located in C:\projects\eval-ad4030_files\2021-07-28-ADI-Kuiper-full-RC-ad4030-v6.img) and choose the directory of the SD card (Go to My Computer and locate the appropriate drive for SD card like E:\). Click on the Flash button and the SD card will start flashing as shown below: -

Figure 15: Win32DiskImager SD card flashing progress

- Now safely eject the USB based SD card writer and disconnect the SD card writer from the PC. Remove the SD card from the SD card writer and insert the SD card into a known good Zedboard™. Ensure that the MIO jumper settings on the Zedboard™ matches with the MIO jumper settings shown in Figure 16.

Figure 16. MIO Jumper (JP7-JP11) Settings on Zedboard

- Now connect the known good EVAL-AD4030-24FMCZ board to the Zedboard in Figure 17.

Figure 17. Connection setup to burn SD cards

Note: This image uses the EVAL-AD4630-24FMCZ, which is the two-channel version of the EVAL-AD4030-24FMCZ

- Connect the micro-USB end of the micro-USB to USB cable to the Zedboard’s USB OTG Port and connect the USB end of the micro-USB to USB cable to the host PC’s USB port as shown in Figure 17.

- Connect the micro-USB end of the micro-USB to USB cable to the Zedboard’s USB UART Port and connect the USB end of the micro-USB to USB cable to the host PC’s USB port as shown in Figure 17.

- Connect the 12V adapter to the +12V connector (J20) on the Zedboard™ and the slide the POWER switch (SW8) on the Zedboard to the ON position as shown in Figure 17. Ensure that the Green POWER LED (LD13) and the Blue Done LED (LD12) on the Zedboard™ turns on as well as shown in Figure 17.

- Wait for the Red LED (LD7) on the Zedboard™ to start blinking. It may take up to 20-25 seconds for the RED LED to start blinking.

- Open Device Manager on the PC and locate the COM port number for the UART terminal. To determine which COM port number is assigned to UART driver when there are multiple UART ports connected to the host PC, un-plug and re-plug the USB cable connected to the UART port of Zedboard™ and note which port disappears from the list and re-appears on in the list.  An example is shown in Figure 18. This port number will be used in the next step.

Figure 18. Finding COM port number in Device manager

- Launch the Tera Term application and in the New Connection window please select the Serial option. Then for the Port choose the Port number of the Serial/UART Port (found in the previous step) from the list of available COM ports. After that go the Setup menu located on top and click on the Serial Port option. Set the speed/baud rate to 115200 under Serial Port setup and connection window. Click on the New Setting button as shown below: -

Figure 19. Tera Term Serial Port Setup

- Enter the iio_info command in the console window  at the TBD prompt to see the IIO devices present on the system and verify if it matches with the output shown in Figure 20. Verify that:

hw_model: field = EVAL-AD4030-24FMCZ on Xilinx Zynq ZED;

iio:device0: field = ad4030-24;

4 channels found

Figure 20: iio-info command executed on Tera Term

- After this turn off the Zedboard™ by sliding the POWER Switch SW8 to OFF position and take out the SD card from the Zedboard™.

- Put the 16GB SD card into one of the SD card plastic case and apply label on the plastic case.

- Take a new empty SD card and program it by repeating the steps from Step 3 to Step 15.

# Appendix 1: Setup SR1 based on EVAL-AD4030-24_SR1_config.xml

Please follow the below steps to setup SR1 signal generator using the EVAL-AD4030-24_SR1_config.xml file:-

- Take an empty USB drive and copy the EVAL-AD4030-24_SR1_config.xml into the drive.

- Turn on the SR1 and connect a USB keyboard and USB drive to the USB ports on top left side and wait for the SR1 to completely turn on.

- Press the windows key on the keyboard. This will open a small window and minimize the main SR1 application

- Remove USB keyboard and connect USB mouse into the USB port and Click on My Computer icon on the opened window. This will open the My Computer folder. Go the D: drive (or the drive where USB-drive files are stored) and copy the EVAL-AD4030-24_SR1_config.xml by right-clicking and selecting copy option.

- Go back to My Computer. Then go to config folder located in C:/Program Files/SR1Audio Analyzer/user/config. Then paste the EVAL-AD4030-24_SR1_config.xml.

- Again open the main SR1 application if it was minimized in the task bar. Once the main SR1 application opens, then click on File menu on top left corner and click on Load Configuration. This will open a standard windows File Open dialog box. Navigate through C:/Program Files/SR1 Audio Analyzer/user/config and select on EVAL-AD4030-24_SR1_config.xml and click open.

- This will open a small window and click OK button and wait for SR1 to configure.

After the configuration is done, please ensure that the Analog Generator window has the following settings: Fs = 512kHz, Mode = Mono, Output = Balanced float with differential impedance = 50Ω, sine wave’s frequency = 1kHz, sine wave’s amplitude =3.5Vrms.