## Slide 1: Test Qualification Plan
SCM Correlation Data Gathering
    Loop 4 bin1 units x30
    Run 100 bin1 units on handler
    Serialize and test 10 bin1 units
    Serialize and test 5 reject units
Ship correlation package from SCM to SCC
SCC Correlation Data Gathering
    Loop 4 bin1 units x30
    Run 100 bin1 units on handler
    Test 10 already serialized bin1
    Test 5 already serialized rejects
SCM/SCC send data to ADGT for Data Crunching and Analysis
CorL8 Analysis of x30 loop /100 units handler data
    X30 loop must pass Mean Shift, Sigma Spread and CPK criteria
    100 Bin1 Correlation units must pass Mean Shift, Sigma Spread and CPK criteria
    10 serialized units must pass bin1 both in SCC and in SCM
    5 serialized rejects must fail the same parameter for both SCC and SCM
Correlation Data Approval
    For TRB movement to Available with Condition
Validation lot run handled by SCC
    Note: CorL8 is ADI data analysis tool.
Reject Correlation |  | 
Unit | SCM | SCC
1 | TnumX: XXXXX | TnumX: XXXXX
… | TnumX: XXXXX | TnumX: XXXXX
5 | TnumX: XXXXX | TnumX: XXXXX
Bin1 Correlation |  | 
Unit | SCM | SCC
1 | Pass | Pass
… | Pass | Pass
10 | Pass | Pass
Correlation Test Criteria(TST00137 ) | 
% Mean Shift Criteria | ( ( SCM_mean - SCC_Mean ) / ( Upper_Limit - Lower_Limit ) ) x 100 < 5
Sigma Spread Criteria | ( SCC_Sigma / SCM_Sigma ) < 1.300000
Cpk Criteria | If CPK to the test limits is >10, then test given automatically PASS

## Slide 2: Test Qualification estimated Timeline
 | PLANNED
 | ACTUAL/ADJUSTED